## Slide 1: Thermal Effects on Critical Flow Venturis
John Wright
NIST Fluid Metrology Group
16th FLOMEKO
September 25, 2013
Paris, France

## Slide 2: Motivation
In a comparison, environmental T sensitivity looks like lab-to-lab differences
If Lab A tests at room T = 23 ºC
and Lab B tests at Room T = 28 ºC?
we can expect a ΔCd
in a 0.4 mm CFV of 0.025%
~ 50 ppm / K
We work at 23.5 ± 1 ºC, but what about our customers?

## Slide 3
Sources of T Sensitivity
4) Sensitivity of reference sensors to room T, e.g. mass flow, pressure
Room T = 297.3 ± 0.5 K
34 and 677 L PVTt in T controlled water bath at 296.7 K
3) Thermal expansion of the throat area
2a = 34 x 10-6 / K, for SS and Cu-Te alloy 145
Significant for large CFVs, less so for small CFVs
d = 30 mm, Jones, Material Temperature Profiles in a Critical Flow Nozzle, ASME, 1983
d = 25 mm, Caron, Kegel, and Britton, 1995, 1996
2) Thermal boundary layer effects: heat transfer from CFV body to gas reduces gas density and mass flux at the CFV throat
1) Temperature “sampling errors” (spatial non-uniformity, stem conduction, time response)
Goal: quantify and correct temperature sensitivity

## Slide 4: Copper (Cu) CFV Design
Four CFV d’s:
3.15, 1.09, 0.648, 0.356 mm

## Slide 5: Experimental Arrangement #1 (to assess T sampling for typical (ISO/ASME) set up)
Bead thermister (TCFV body)
Heater controller RTD
Tz (“hot wire” T sensor)
Ceramic pipe liner
Tup1
Tup3
Tup2
TCFV body
Cu CFV
Water to gas heat exchanger
Water pump

## Slide 6: Cd Measurements
200, 300, 400, 500, 600, and 700 kPa
3 PVTt collections at each pressure
On 2 occasions, average of 6 points
U(Cd) = 0.06 %

## Slide 7: Spatial Temperature Variations (d = 3.15 mm, no CFV heating)
Tup1
Tup3
Tup2
TCFV body
Tz

## Slide 8: d = 3.15 mm, no CFV heating
0.05 %
Tup3 is cold due to conduction from CFV body through approach pipe wall
Tup1
Tup3
Tup2
TCFV body
Tz

## Slide 9: d = 3.15 mm, TCFV body = 313 K
Tup1
Tup3
Tup2
TCFV body
Tz
Add a PID controlled heater to CFV: 298, 303, 308, 313 K

## Slide 10: d = 3.15 mm, TCFV body = 313 K
0.39 %
Tup3 is hot due to conduction from CFV body through approach pipe wall
Tup1
Tup3
Tup2
TCFV body
Tz

## Slide 11: Experimental Arrangement #2 (to minimize T sampling uncertainties)
Water to gas heat exchanger
Water pump
Optional room T water jacket
Optional heater

## Slide 12: d = 3.15 mm, TCFV body = 313 K
Tup1
Tup3
Tup2
TCFV body
Tz
0.01 %
All 4 T sensors give the same Cd values within 0.01 %
Add a 297 K water bath to the approach pipe

## Slide 13
Numerous Heat Transfer Mechanisms
Measured T0 is subject to sampling errors due to heat transfer within the flow, CFV body, and approach pipe walls
Internal flow
CFV body
Room
Approach pipe wall
Inlet gas
T sensor

## Slide 14
Assume 1-D, isentropic, adiabatic flow through a ISO toroidal copper CFV, 2.2 cm body radius, Bartz 1965 convective heat transfer coefficient, no axial direction heat transfer…
What is the temperature distribution in the gas and CFV body?
d = 3.15 mm
d = 0.36 mm
Tcore
Tadiabatic wall
TCFV ext
TCFV int

## Slide 15: T Distribution in CFV Body
By using high thermal conductivity CFV material, TCFV body is close to inner wall T
(Note that low thermal conductivity material is desirable for non-research applications!)
Material | k
 | [W/(cm K)]
Cu | 3.8
SS | 0.16
Macor | 0.0146
Kegel and Caron, ASME Fluids Engineering Summer Meeting, San Diego, California, USA, 1996.

## Slide 16: CT Measurements
Four Cu CFVs: d = 3.15, 1.09, 0.65, and 0.36 mm
Use analytical Cd values and TCFV body= 298 K data to determine d
PID control of TCFV body = 298, 303, 308, and 313 K
Apply 2aDT thermal expansion corrections (Tref = 298 K) and plot differences in Cd relative to values at 298 K
Assume CT = 1 - DCd (because other known effects have been corrected)
298 K
303 K
308 K
313 K
298 K
303 K
308 K
313 K

## Slide 17
Bejan, Heat Transfer, 1993
There is a velocity boundary layer AND a thermal boundary layer
The warmer, lower density layer near wall leads to lower flow than adiabatic assumption
Johnson, A. N., 2000
CFV theoretical mass flow equation assumes adiabatic wall (no heat transfer from CFV body to gas)
In reality, a thermal boundary layer is present

## Slide 18
298 K
303 K
308 K
313 K
A Correction for the Laminar Thermal Boundary Layer
Choose a reference TCFV body (298 K)
Assume CT = 1 - DCd (because other known effects have been corrected)
As for Cvbl, CT scales with Re-1/2
CT is proportional to the density change in the thermal boundary layer relative to some reference condition, i.e. proportional to (Tref -TCFV body) / Tref
K = empirical constant

## Slide 19
A Correction for the Laminar Thermal Boundary Layer
Laminar to turbulent boundary layer transition leads to a Reynolds number offset (Re0)

## Slide 20
In prior studies:
Bignell, N. and Choi, Y. M., Thermal Effects in Small Sonic Nozzles, Flow Meas. Instrum., 13, pp. 17 – 22, 2002
2.04 mm
0.715 mm
1.38 mm
1.0 mm

## Slide 21
0.36 mm CFV, T Sampling?

## Slide 22
Because of CFV body cooling, when we perform a “Room T” calibration, we are following an arc in the CT plane…
298 K
303 K
308 K
313 K
288 K ?
293 K ?

## Slide 23
Conclusions
In small CFVs (< 10 mm?), temperature sampling errors and thermal boundary layers lead to significant temperature sensitivity (50 to 300 ppm/K)
Making thermal expansion corrections in small CFVs makes the temperature sensitivity worse!
There are complex heat transfer mechanisms and significant temperature gradients in CFV installations
Better designs of approach pipe, T sensors, sensor placement, and CFV materials will reduce CFV calibration reproducibility and sensitivity to room T
A simple physical model for the thermal boundary layer matches experimental CT values, i.e. corrections are possible